## Tools
Tool | Reader study | Panel readers | Examinations | Readers' average area, unaided | Readers' average area, with the tool | Gain | 95% interval, lower | 95% interval, upper | Meets the protocol | Readers who did better with the tool | Why not
Lumidra Detect | MBS-AI-01 | 10 | 300 (75 cancers, 225 normal) | 0.8179 | 0.8651 | 0.047 | -0.015 | 0.109 | No | 10 | The gain varies from examination to examination, the same way for every reader, so it is not shown for the service's screening examinations.
Ondrix Mammo | MBS-AI-02 | 10 | 300 (75 cancers, 225 normal) | 0.8607 | 0.9027 | 0.042 | -0.013 | 0.097 | No | 7 | The gain differs widely from reader to reader (3 of 10 read worse with it), so it is not shown for the service's film readers.
Halvic Breast | MBS-AI-03 | 10 | 300 (75 cancers, 225 normal) | 0.8586 | 0.8846 | 0.026 | 0.014 | 0.038 | Yes | 9

## Readers
Tool | ReadPoint user | Reader | Role | Area unaided | Area with the tool | Change
Lumidra Detect | u10474 | FR13 | Consultant radiographer | 0.869 | 0.875 | 0.006
Lumidra Detect | u17234 | FR26 | Consultant radiographer | 0.802 | 0.841 | 0.038
Lumidra Detect | u19943 | FR15 | Breast clinician | 0.784 | 0.855 | 0.072
Lumidra Detect | u28756 | FR21 | Consultant radiologist | 0.827 | 0.856 | 0.029
Lumidra Detect | u32643 | FR03 | Breast clinician | 0.862 | 0.871 | 0.008
Lumidra Detect | u36906 | FR01 | Consultant radiologist | 0.773 | 0.866 | 0.093
Lumidra Detect | u52000 | FR16 | Consultant radiologist | 0.821 | 0.866 | 0.044
Lumidra Detect | u60139 | FR02 | Advanced practitioner radiographer | 0.852 | 0.903 | 0.05
Lumidra Detect | u61835 | FR17 | Consultant radiographer | 0.781 | 0.85 | 0.069
Lumidra Detect | u95756 | FR24 | Consultant radiographer | 0.807 | 0.87 | 0.063
Ondrix Mammo | u33675 | FR01 | Consultant radiologist | 0.859 | 0.85 | -0.009
Ondrix Mammo | u39432 | FR02 | Advanced practitioner radiographer | 0.895 | 0.929 | 0.034
Ondrix Mammo | u44845 | FR09 | Breast clinician | 0.846 | 0.781 | -0.065
Ondrix Mammo | u45192 | FR12 | Breast clinician | 0.809 | 0.914 | 0.105
Ondrix Mammo | u48068 | FR08 | Advanced practitioner radiographer | 0.913 | 0.923 | 0.01
Ondrix Mammo | u50012 | FR21 | Consultant radiologist | 0.895 | 0.96 | 0.065
Ondrix Mammo | u52812 | FR19 | Advanced practitioner radiographer | 0.857 | 0.807 | -0.05
Ondrix Mammo | u62416 | FR11 | Consultant radiologist | 0.842 | 0.967 | 0.125
Ondrix Mammo | u77851 | FR17 | Consultant radiographer | 0.853 | 0.944 | 0.091
Ondrix Mammo | u93344 | FR14 | Advanced practitioner radiographer | 0.838 | 0.952 | 0.114
Halvic Breast | R1245 | FR14 | Advanced practitioner radiographer | 0.877 | 0.912 | 0.035
Halvic Breast | R1508 | FR22 | Advanced practitioner radiographer | 0.915 | 0.935 | 0.019
Halvic Breast | R2694 | FR15 | Breast clinician | 0.843 | 0.848 | 0.004
Halvic Breast | R3309 | FR05 | Consultant radiologist | 0.807 | 0.83 | 0.023
Halvic Breast | R3871 | FR03 | Breast clinician | 0.858 | 0.891 | 0.033
Halvic Breast | R4785 | FR17 | Consultant radiographer | 0.838 | 0.88 | 0.042
Halvic Breast | R5036 | FR04 | Consultant radiologist | 0.888 | 0.926 | 0.038
Halvic Breast | R7820 | FR07 | Consultant radiologist | 0.859 | 0.894 | 0.035
Halvic Breast | R9545 | FR18 | Advanced practitioner radiographer | 0.8 | 0.835 | 0.035
Halvic Breast | R9659 | FR08 | Advanced practitioner radiographer | 0.9 | 0.895 | -0.005

## Sites
Site | Screening studies sent for reading, 1 Oct 2025 to 30 Sep 2026 | Of which from mobile units
Corrowby Road Breast Unit | 11569 | 5152
Wendlow Breast Screening Centre | 7298 | 2096
Copse Lane Breast Screening Centre | 13253 | 7472
Tallowmere Park Breast Unit | 13758 | 7615
All sites | 45878 | 22335

## Costs
Tool | Charged on | Volume | Unit charge (GBP) | Annual charges (GBP) | One-off (GBP) | First contract year, excluding VAT (GBP)
Ondrix Mammo | screening studies received for analysis | 45878 | 1.62 | 74322.36 | 14800 | 89122
Lumidra Detect | screening episodes; film reader accounts | 44,262 episodes; 27 accounts | 2.08 per episode; 340 per account | 101244.96 | 9500 | 110745
Halvic Breast | film reader seats; screening sites | 27 seats; 4 sites | 2,450 per seat; 5,900 per site | 89750 | 12000 | 101750

## Film readers
Staff group | Film readers in post on 1 April 2027
Advanced practitioner radiographer | 9
Breast clinician | 4
Consultant radiographer | 6
Consultant radiologist | 8
All film readers | 27

## Notes
Measure | Each reader's area under the empirical ROC curve of their probability-of-malignancy ratings against the study set's reference standard (ties counted half), the reading of record being the reader's last submission; the gain is the panel's average area with the tool minus its average area unaided (protocol 3).
Interval | Two-sided 95 per cent, with the panel's readers and the study's examinations both treated as samples of the service's film readers and screening examinations (protocol 4.1): Obuchowski-Rockette analysis with jackknife covariances and Hillis's degrees of freedom, the same as DBM; DeLong covariances and a bootstrap of readers and examinations give the same verdicts.
Studies | Completed routine and technical repeat screening studies dated 1 October 2025 to 30 September 2026. Records the legacy system's history load added for studies already registered in the new system during each site's parallel running are counted once; implant views registered as a second study are counted. A mobile unit's studies count to the site hosting it on the day (mobile visit schedule): unit 2's earlier vehicle to Corrowby Road until 29 May 2026.
Episodes | One per woman and screening round, counted with its routine study (Lumidra's quote): technical repeats, including a later-day recall booked under the routine code, belong to it, and a repeat in early October 2025 belongs to an episode that began before the 12 months.
Film readers | Staff who read films, in post on 1 April 2027 (start on or before, end on or after), counted once per professional registration number (the bank system records the number without the GMC or HCPC prefix): one reader holds a substantive and a bank contract.
Prepared | AI evaluation lead, 10 November 2026.